## Slide 1: FDA Critical Path Project
“Generation and Review of Regulatory-Grade Computational Evidence Using a Generic Inferior Vena Cava Filter”
Summary of Preliminary FEA

## Slide 2: The objectives of this project are to…
???
Develop a generic IVC filter
Predict & demonstrate:
  fatigue resistance
  embolus-trapping efficiency
  Ideally,
  single strut shape (simplify fatigue simulation/testing)
  resist tilt, migration, and vein wall perforation
*designs investigated in previous FEA simulations
Notes: Our objective is to develop a “generic” IVC filter for which we will generate “regulatory grade” computational evidence.
Specifically, we will predict & demonstrate the fatigue performance and embolus-trapping efficiency of the device.

## Slide 3: FEA performed to evaluate filter design
breathing
Valsalva
Pelton, A. R. (2011). J. Mater. Eng. Perform., 20(4-5), 613–617
?
simulated CI = 0.5
(see Laborda et al. 2014)
collapsibility index
CI = (dmax- dmin)/dmax
IVC
cross-
section
dmin
dmax
Setup:
ABAQUS Dynamic/Implicit (quasi-static); 10x mass scaling
C3D8I elements (4 per strut thickness & width; ~3,000 to 6,000 per strut)
  SE508 nitinol (material properties from Craig Bonsignore’s example on GitHub)
Simulated conditions → extracted quantities
Sheathing → prestrain
Filter placement → contact force/area
Valsalva → mean/amplitude strain
Two IVC diameters considered (human variability):
14mm (3.5mm wall displacement during Valsalva)
28mm (7.0mm wall displacement during Valsalva)
We assume Valsalva is the limiting loading scenario for fatigue life
Notes: FEA simulations were used to evaluate the performance of the candidate designs
A mesh refinement study was also performed (4x4 versus 8x8 elements per strut cross-section; difference in maximum strain amplitude of 10%)
Valsalva is the limiting case for fatigue life (larger displacements than those associated with respiration)
N ≈ 4 million (“worst case” scenario assuming chronic cough for 10 years; see discussion in CP project outline & Birring et al. 2006)
CI = 0.5 +/- 0.04 , n=100 patients (see discussion in CP project outline and Laborda et al. 2014)
Strain amplitude and mean strain are calculated using a tensor-based approach, e.g.: E_strain_amplitude_scalar = max_abs_principal{0.5*(E_logarithmic_strain_tensor_loaded – E_logarithmic_strain_tensor_unloaded)}
Further justification needed for chosen “performance measurements” and their target values:
“Contact length” instead of “contact area”? (quantification of contact area with FEA is challenging; accurate results require very fine meshes. However, the predicted trends may still be useful)
What is a reasonable range for the per-strut “contact force”? (see Dowell et al. 2015, Garcia et al 2012; vein wall perforation observed around 1N (upper limit))
Are there other quantities we should consider?

## Slide 4: Performance goals are set for quantities extracted from FEA simulations
 | distal end contacts wall? | prestrain (max/min prin. LE) | contact force (N) (per strut) | contact area (mm2) (per strut) | ε amplitude, Valsalva (CI=0.5) | mean ε, Valsalva (CI=0.5)
target | yes | <6% | 0.010 to 0.100 | >0.05? | <0.4% | <6%
*to avoid plastic deformation during sheathing
Notes: FEA simulations were used to evaluate the performance of the candidate designs
A mesh refinement study was also performed (4x4 versus 8x8 elements per strut cross-section; difference in maximum strain amplitude of 10%)
Valsalva is the limiting case for fatigue life (larger displacements than those associated with respiration)
N ≈ 4 million (“worst case” scenario assuming chronic cough for 10 years; see discussion in CP project outline & Birring et al. 2006)
CI = 0.5 +/- 0.04 , n=100 patients (see discussion in CP project outline and Laborda et al. 2014)
Strain amplitude and mean strain are calculated using a tensor-based approach, e.g.: E_strain_amplitude_scalar = max_abs_principal{0.5*(E_logarithmic_strain_tensor_loaded – E_logarithmic_strain_tensor_unloaded)}
Further justification needed for chosen “performance measurements” and their target values:
“Contact length” instead of “contact area”? (quantification of contact area with FEA is challenging; accurate results require very fine meshes. However, the predicted trends may still be useful)
What is a reasonable range for the per-strut “contact force”? (see Dowell et al. 2015, Garcia et al 2012; vein wall perforation observed around 1N (upper limit))
Are there other quantities we should consider?

## Slide 5: Stress-strain history diagram (Rev1)
εpre-strain
2 x εamp
εmean

## Slide 6
1DOF
contact with friction
*FEA simulations of Valsalva shown for reference; note that the contact point between the IVC wall and the filter strut changes as the strut deforms
IVC wall
contact point(s)
7mm
3.5mm
28mm
IVC diameter
14mm
Valsalva (fatigue loading) (Rev1)

## Slide 7: Results summary (Rev1)
28mm
IVC diameter
14mm
IVC diamter | distal end contacts wall? | prestrain (max/min prin. LE) | contact force (N) (per strut) | contact area (mm2) (per strut) | ε amplitude, Valsalva (CI=0.5) | mean ε, Valsalva (CI=0.5)
14mm | yes | 0.80% | 0.028 | 0.018 | 0.07% | 0.45%
28mm | yes | 0.80% | 0.010 | 0.018 | 0.12% | 0.24%
Valsalva
*results from third Valsalva cycle
Filter Placement
Strain Limit Diagram (SLD)
IVC wall
FS ≈ 3.33?

## Slide 8: Mesh refinement (Rev1)
IVC diamter | distal end contacts wall? | prestrain (max/min prin. LE) | contact force (N) (per strut) | contact area (mm2) (per strut) | ε amplitude, Valsalva (CI=0.5) | mean ε, Valsalva (CI=0.5)
coarse (4x4) | yes | 0.80% | 0.0096 | 0.018 | 0.121% | 0.24%
medium (8x8) | yes | 0.92% | 0.0096 | 0.0048 | 0.134% | 0.26%
fine (12x12) | yes | 1.04% | 0.0102 | 0.0050 | 0.139% | 0.27%
GCI: -2.81%
coarse
medium
fine

## Slide 9: Eccentric loading of filter strut (Rev1)
IVC diamter | distal end contacts wall? | prestrain (max/min prin. LE) | contact force (N) (per strut) | contact area (mm2) (per strut) | ε amplitude, Valsalva (CI=0.5) | mean ε, Valsalva (CI=0.5)
28mm | yes | 0.80% | 0.010 | 0.018 | 0.12% | 0.24%
28mm | yes | 0.80% | 0.0086 | 0.0091 | 0.14% | 0.26%
a
b
a
b

## Slide 10: Static/Implicit vs. Dynamic/Implicit (Rev1)
IVC diamter | distal end contacts wall? | prestrain (max/min prin. LE) | contact force (N) (per strut) | contact area (mm2) (per strut) | ε amplitude, Valsalva (CI=0.5) | mean ε, Valsalva (CI=0.5)
28mm | yes | 0.798% | 0.00884 | 0.0182 | 0.122% | 0.239%
28mm | yes | 0.797% | 0.00959 | 0.0182 | 0.121% | 0.238%
Static/Imp
Dynamic/Imp
*Simulations were performed using Abaqus Standard, Dynamic/Implicit to increase stability of contact interactions. Simulations performed using Static/Implicit yield similar results.
Dynamic/Imp